## Slide 1: Cathode production and supply Photocathode website (useful things to track)
Daniele Sertore INFN Milano – LASA

## Slide 2: Introduction
In the last 15 years we have delivered about 120 cathodes mainly to DESY-FLASH and DESY-PITZ
We have a mean operative lifetime of the cathode (24/24h 7/7d) of about 3 months with FLASH parameter (1 nC, 800 ms, 10 Hz, 42 MV/m)
Mean QE at delivery is about 10 % (l=254 nm)

## Slide 3: Cs2Te GROWTH and diagnostic

## Slide 4: Cs2Te Recipe
A. di Bona et al., JAP80(1996)3024
 | Cs/Te
b) | 0.6
c) | 1.2
d) | 1.9
e) | 2.5
Cathode growth
The recipe was developed in Milano and as a “process” transfer to Modena for analysis.
Auger Electron Spectroscopy shows formation of different stoichiometric compounds during Cs deposition. The “correct” Cs/Te ratio is reached corresponding to the QE maximum.
Substrate contribution
The Mo plug prevents diffusion of the film (AES + Ar+ sputtering study). Estimated photocathode typical thickness is 20-30 nm.
Cs2Te
Mo
Cs2Te + Cu
Cu

## Slide 5: The Cs2Te support
The recipe was then transfer to a plug compatible with gun operation
The Mo plugs, either from sintered or arc-cast material, are cleaned with a Buffered Chemical Polishing after machining.
The plugs are then polished to optical finishing to reduce dark current
The quality of the optical finishing is measured by photocathode reflectivity which provides also information on surface roughness, typically  10 nm.
Sintered
Rough Surface
Optical Finished
Arc-cast

## Slide 6: Deposition and Diagnostic
During cathode deposition a circular masking system shapes the photoemissive layer and center it. The actual masking diameter is 5 mm
5 mm
Active Area
Hg lamp
Filter wheel
Shutter
Remote controller
Photocurrent signal
Prep. chamber viewport
Photodiode
Reflected power
Incident power
Calibrated photodiode for reflected light power from the film
A motorized and remotely controlled filter wheel with interference filters (239 nm, 254 nm, 334 nm, 365 nm, 405 nm, 436 nm)
Picoammeter for photocurrent measurement

## Slide 7: On-line diagnostic
Control
Te deposition on the Mo plug
Completion of the Cs2Te (no Cs excess)
  Measure
QE & R @ ls during deposition
  Estimate
Eg+Ea threshold formation during the deposition
R @ 254nm vs. Te and Cs deposition
QE @ 254 nm vs. Cs deposition
The “last” peak
QE @ different l during the Cs deposition
5 nm
10 nm
15 nm
Eg+Ea and QE @ 254nm vs. Cs deposition
QE
Eg+Ea

## Slide 8
  The spectral response analysis provides the energy threshold (Eg+Ea).
  We fit data with function based on the “Kane” model
  Typical values for thresholds:
    3.3 eV (“high energy”)
      2.8 eV (“shoulder”)
Spectral Response and QE maps
Eg+Ea mapping
Fresh Cathode
Aged Cathode
QE [%]
mm
mm

## Slide 9: Spectral response reproducibility
The on-line diagnostic allows better control of the cathode growing.
The reproducibility of produced cathode spectral responses is largely improved.
The Cs excess is under control. No more low energy shoulder.
104.4
417.1
77.3
123.2
107.3

## Slide 10: Post Usage Diagnostic
Cathode 13.4 – 157 days of operation
Measurement done at LASA
QE @ 254 nm = 3.5%
Initial QE was 11%.
Decrease of the low energy threshold as expected
Cathode 77.2 – 176 days of operation
After usage, central part of the cathode has a clear darker area
Lack of coating on the right upper film area
After production
After usage

## Slide 11: R&D activities
Optical Properties
Thermal Emittance
Rs
Rp
n | k | s [nm] | thickness [nm]
1.75 | 0.356 | 0 | 38.5
Ag
Cs2Te

## Slide 12: PHOTOCATHODE SYSTEMS

## Slide 13: INFN Photocathode Systems
Preparation Systems | Transfer System
INFN Milano – LASA | DESY - FLASH
DESY – HH | DESY - PITZ
Fermilab – Lab7 | DESY - REGAE
Fermilab-A0 | Fermilab –NML
HZB (XPS) | LBNL

## Slide 14: Cathode systems
LASA preparation system
Preparation Chamber @ FNAL
RF Gun and FLASH linac
Transfer Chamber
FNAL Transfer system

## Slide 15: Carriage Evolution
A0 Carrier
2nd gen.
Bajonet
Increased Space
New Sliding Blocks
3rd gen.
4th gen.

## Slide 16: Transport System
Cathodes are transported under UHV condition from INFN Milano to the Labs.
Transport box
Cathode Loading
SAES D100 NEG Pump
No power supply needed. OK for airplane
Ion Getter Pump
Triax C.C.

## Slide 17: Cathode 409.1 in LBNL Box
The efficiency of the new pumping system, has been tested:
switching off the ion pump (for about 3 months)
measuring the QE vs. ls
QE at 254 nm:
  Stable within few percents
  QE at 436 nm:
  Reduced by an order of magnitude
High stability of the spectral responses
1 Oct 2012

## Slide 18: Photocathode handling system
Analysis of accident during cathode operation:
  Only three significant accidents have occurred after about 12 years of operation
Damages on the connection system:
  due to a not correct parking of the carrier in the transport box and then closing of the valve
  due to the manipulator magnet uncoupling from its vacuum counterpart
12

## Slide 19: Photocathode database

## Slide 20: The Photocathode Database
The database tracks photocathode performances in the different transport boxes and in the different labs as well as properties of plugs during cathode preparation.
Data from INFN Milano, DESY (FLASH, PITZ, REGAE), FNAL (NML & A0) and LBNL are included.
Box Driven
Plug Driven
Access to plug data and cathode related properties
Access to cathode properties during operation

## Slide 21: Database filling
MS Access Forms
Dedicated program
Web Interface

## Slide 22: Web Interface
http://wwwlasa.mi.infn.it/ttfcathodes/

## Slide 23: From Rod to …
Reflectivity
Lapping
BCP
Heating
Loading
Deposition

## Slide 24: … Photocathode

## Slide 25: Some Statistics

## Slide 26: Cs2Te QE after production
QE[%] = 10.2±2.1%

## Slide 27: QE old and new diagnostic
QE[%] = 8.6±2.7%
QE[%] = 11.2±1.6%

## Slide 28: QE during operation
QE during operation.
The operation last for more than 100 days of continuous operation.
The cathode is changed if:
QE < 0.5 %
dark current is limiting LINAC operation

## Slide 29: Dark current and Guns
Trend of dark current for different guns. To be notice DC decrease during operation of the gun.

## Slide 30: Lab Overviews

## Slide 31: Future work
Include “spectral responses” (Eg+Ea)
Include “QE map” (uniformity)
Include deposition parameters:
  Te thickness (as representative of the deposition process)
Include a sort of generic “query” for retrieving data from web interface

## Slide 32
THANK YOU
A special acknowledgement to
Laura Monaco and Paolo Michelato (INFN)
Sven Lederer and Siegfried Schreiber (DESY)