## Slide 1: SRF Cavities Resonance Control. CW mode of operation (FNAL’s experience).
Yu. Pischalnikov W. Schappert
FNAL
TTC CW SRF Meeting, Cornell University, 12June, 2013

## Slide 2
Microphonics for CW SRF machine
  dF/dP: cryo-plant OR/AND cavity system (cavity/HeVessel/tuner) design (see T. Khabiboulline – FNAL’s design of 650MHz &325MHz design);
  External vibration: design of cryo-module & vacuum/cryo infrastructure around machine;
  Active compensation of the microphonics /challenges/:
    Accuracy of detuning measurements
    Slow feed-back compensation (slow dF/dP compensation )
    Fast feed-back compensation (external vibration; mechanical model of cavity)
    Long Term stability of active compensation (one trip/per cavity/500Hours operation)
Tuners coarse(slow)& fine(fast)
  Design & Reliabilities (piezo reliability)

## Slide 3: Active control of SRF cavities resonance
 | operation mode | 
algorithm | CW | RF pulse
Feed-forward |  | 
Slow Feedback (He pressure fluctuation & cavities dF/dP) |  | 
Fast Feedback (vibrations) |  | 

## Slide 4: Adaptive Feed-forward algorithm for LFD compensation for 1-10ms (and longer) pulse.
Accurate measurements of the detuning due to Lorentz force
Accurate characterization of the detuning response to piezo impulses at various delays
Use least-squares to determine the combination of piezo impulses required to cancel out the Lorentz force detuning
Lorentz Force Detuning Before and After Compensation for one 9-cell cavity equipped with blade tuner
10ms RF pulse
1ms RF pulse
Flat top

## Slide 5: Adaptive Feed-forward @RF pulse mode CM1 NML (He pressure & dF/dp)
Frequency stability of the CM1 cavities over the course of 12 hours with non-adaptive LFD compensation. The cavity resonance frequencies can drift by several tens of Hz.
20 hours
Fermilab CM1 tuning stability with adaptive feed-forward LFD compensation ON. CM1 operated at 5Hz repetition rate. Adaptive feed forward (adjusting piezo bias for next RF pulse) can limit the drift to ~1 Hz.

## Slide 6: Slow Feed-back to compensate microphonics (as a results of He pressure fluctuation ( dP~5torr @4.5k) and cavity’s dF/P ~140Hz/torr )
s=0.45Hz
s=1.3Hz
Half Bandwith:1-2Hz
Active Compensation of Microphonics in the SSR1 Cavity. The green curve shows the pressure of the helium bath. The blue curve shows cavity detuning. The cavity was operating with a bandwidth of ~1.5Hz at the temperature 4.5K and Eacc=5MV/m.
FPGA fed Df (detuning) back to fast tuner.
Simple simulations indicated that limit may be 100kHz ADC
Upgraded hardware: Used ADC14-bit, 100MS/s
SSR1 cavity was operated with bandwidth of ~102Hz
Goal was to see ultimate level of compensation

## Slide 7: Fast Feed-back @CW
CC2 (analog phase detector) & filter bank
Single Narrow Resonance (F=92Hz; T=2K)
Active damping reduces the vibration at this frequency
by 15 dB.
Frequency(Hz)
Time (s)
Amplitude (FFT)
Time (s)
The Piezo tuner actively damped vibrations induced by external sources. An IIR filter bank was used to isolate and reverse the phase of the particular spectral line. The phase reversed signal then fed back to the piezo tuner.
9cell 1.3GHz “tesla style cavity (CC2)
Now is part of FNAL’s NML photo-injector

## Slide 8
B
Directional coupler
Detuning calculation
Pprobe
Pforw
Prefl
LLRF
Df(detuning)
Piezo
Piezo stimulus
FPGA
/algorithm
/Piezo_amp
Cavity mech.model
Calculate detuning from RF signals using the same detuning formula used for LFD compensation
  Directional Coupler issues: Contamination of Pforw with Prefl
  Quality of RF receivers and digital filtering contamination of our baseband signal
Cavity’s mechanical model to feed into FPGA system
Double structure distribution… 0.5Hz resolution instead of 0.18Hz
Assumption (likely only partially true) that forward tail is due exclusively to cross-contamination by reflected signal.
Correction reduces the amplitude and phase variation and forward signal from 5-15% to a few percent.
1(b)
1(a)
Correction used at FNAL for RF-pulse
mode LFD compensation
Plans at FNAL related to microphonics work
Corrected
Uncorrected
phase
magnitude

## Slide 9: Measuring the Mechanical Transfer Function in Pulsed-Mode (opportunity to do it at S1-Global)
Excite the piezo with a series of impulses and sweep the impulse-to-RF delay
  Equivalent to measuring the piezo-to-detuning transfer function using CW
Delay-Scan used to determine optimal waveform
KEK cavity
Blade/FNAL
DESY/Saclay

## Slide 10: CM1 Cavities (#5&#6 of NML cryomodule#1) measurements and 20-Pole MSSR Model

## Slide 11: Summary
FNAL focus is changed from RF-pulse machine (ILC) on CW machine (Project X)
Development of expertise in active compensation of the mircophonics is important goal for FNAL’s team
We are interesting to collaborate with experts in SRF community:
  exchanges ideas; software codes; joint work at existing facilities; tuner’s reliabilities study; etc…

## Slide 12: Additional Slides

## Slide 13: Fast Feed-back @CW
Single Wide Resonance (F=155Hz; T=2K)

## Slide 14: Measuring LFD
Complex envelope of the probe signal described by 1st order ODE
Rearrange terms to extract the width and the dynamic detuning from the probe and forward IF signals

## Slide 15: Long Term Stability