## Slide 1: SRF photo-injector and Booster modules at bERLinPro: Assembly and commissioning status
Can progress be measured?
10/04/22, A. Neumann for the Sealab/bERLinPro SRF Team*
66th ICFA Advanced Beam Dynamics Workshop on Energy Recovery Linacs
Cornell University, Ithaca, NY
*A. Frahm, F. Göbel, S. Klauke, H. Plötz, M. Schuster, Y. Tamashevich, J. Ullrich

## Slide 2: Challenges for the SRF+vacuum systems
Introduction
SRF Gun:
Low emittance high brightness beam
Insertion of a thermally isolated high QE Cs-K-Sn cathode (particulate free)
Control of field emission to avoid dark current based losses and halo formation
High current, high beam power operation (HOM, FPCs)
High emission RF field  beam dynamics optimized shape
Avoid coupler kick-based emittance dilution
Synchronization Laser  RF, precise RF control
2.3 MeV, 100 mA
High beam power
High beam current
Intermediate field levels 10 MV/m
Low coupler kicks
Zero-crossing operation
SRF booster cavity (Cornell)
+4.5 MeV (100mA)
Low beam power
High beam current
Strong higher order mode damping
Higher field levels: 19 MV/m
Multi-pass beam
Precise field control and tuning
Main linac cavity
+45 MeV (2x 100 mA, recovered)
Compact machine: All vacuum parts need to be treated within ISO 4-5 environment
Complete machine analyzed and optimized by wake/impedance studies
Movable local cleanrooms ISO 5, rail system on the ceiling for proper placement

## Slide 3: Status of the SRF gun
SRF gun
2xCW modified TTF-III
Coupler: Qext 3.6.106
for up to Iavg=6-10 mA,
10 kW each
Study 2 coupler operation
106 mm beam tube:
Allows propagation of
lowest TM110 mode:
HOM studies
Blade tuner with
motor and 4x piezo tuner:
Microphonics compensation
3 pick-up antennas to measure HOM
polarization
0.4.l/2 cell + full cell:
Optimized
emission phase
HZDR cathode
insert and choke
cell design:
Proven system
Cathode exchange
with HZDR
Stiffening ring:
Df/DP minimized
Chimney 22 cm²~35 W at 1.8 K
about Epeak=45 MV/m at Q0=3.5.109

## Slide 4: Evolution of the SRF Gun at HZB
SRF photo-injector: From cradle to grave and resurrection
Time
2013
VTA
Manufacture
2014-2016
HZB
VTA/HTA
Test/Assembly
HTA
String test
HZB
Module assembly
Gunlab
String/Cold mass
2016
2017/2018
Commissioning
Cathode transfer
1st Beam
2018-2021
2011-2012
HZB
Design
PC
Fully in-house developed SRF photoinjector with strong partners
SRF gun history
ERL17
1.1
1.0
Damage and Recovery
2022
Reassembly
ERL19

## Slide 5: Cold String layout of bERLinPro SRF Gun
Parameters
SC solenoid
Twin CW-modified
TTF-III coupler up to 20 kW
x, y steerer
SiC HOM
absorber (Cornell)
Cathode tube with tuner (HZDR)
Gate valve
and transition
to 300K
Cavity+LHe vessel (HZB, JLab),
Cornell-type blade tuner,
2 layer magnetic shielding
Support
structure
*
*
*
80K
80K
4K
1.8K
4K
1.4xl/2 cells + Choke cell
*15-20% higher
found by beam
energy measurements
Prototype designed by HZB
manufactured by JLab

## Slide 6: Operation of first 1.4 cell SRF Gun with cathode exchange mechanism
First beam
First beam 12/2017
SRF Gun module
01/2018 Operation with high QE cathode failed,
as cathode plug dropped into Gun cavity by malfunction of holder mechanism 

## Slide 7: Development of high QE cathode and drive Laser
High QE cathode
doi:10.1103/PhysRevAccelBeams.21.113401
Last evidence of cathode being alive
Photocathode preparation and analysis laboratory up and running since 2015
- insight into growth process with material science studies in parallel
- achieved quantum efficiency of 16.8% at 515 nm (specs 1%)
- successful transfer from cathode lab to SRF gun
- demonstrated, that cooling of cathode to 120 K do not harm quantum efficiency (unlike prediction by other researchers!)
We were close to
operate this record cathode in Gunlab…….
And learned from our
fails…..
…cathode holder
heating and shock experiments, Na based cathodes…..and the cavity?

## Slide 8: Recent Achievements  Refurbish SRF Gun
repair
1.1
1.0
Y. Tamashevich, et al. published SRF21
Both cavities were damaged: 1.0 Initial scratch after production  Dark current 1.1 got damaged at manufacturer
A repair and refurbishment program was set up between 06/2019 until 06/2021
Treatment by grinding, polishing, BCP and optimized HPR system, HF rinsing 
1.1: HPR lance scratch
1.1 After grinding and polishing
1.0: surface defects
Gun 1.0 is back on track
Gun 1.1 close to old level
Gun 1.0

## Slide 9: A closer look to the procedure and testing
Repair
Prototype cavity served as pilot to test and develop all procedures:
Defect analysis by replica method and laser microscopy
Defect grinding by special polishing method in high E-field area
BCP+HPR by vendor, later standard in-house HPR
In-house new HPR nozzle head for SRF gun geometry (several iteration)
HF rinsing of oxidization and following HPR with found parameters
Y. Tamashevich et al., Superconductor Science and Technology
To be published
1.0
1.1
F. Göbel

## Slide 10: Dark current analysis of Gun 1.0 Could we have continued with Gun 1.0 as it was from the start?
Dark current
Measured dark current distributions with Solenoid scan (left) on 1st screen
and Idark (E0) at different stages of assembly/operation
Simulated dark current on 1st view screen
Goal:
Recreate situation of Gun 1.0 at Gunlab and assembly stages
Simulate dark current from the found defects close to the cathode hole opening
Analyze dark current distribution in injector by simulation and varying parameter towards ERL range
Simulated propagation
Upstream
1st viewscreen
Adjust Fowler Nordheim field emission to measured
values.
Measured dark current downstream is only fraction of real amount of current and depends on solenoid settings
E0=30 MV/m
100µA vs 100 mA at cavity exit
20µA vs 100 mA 1m downstream: 2.10-4 losses!!!
1st viewscreen
Cavity exit
2nd viewscreen

## Slide 11: Screen based beam measurements with dark current
Gunlab and defect analyzed
Ekin (eV)
Screen size
Beam tube diameter
Beam distribution with dark current by two backwall defects and E0=30 MV/m Scan of 20 ns, aka laser rep. rate of 50 MHz, afterglow effect of scintillating screen?
77 pC beam charge vs. 0.5 pC dark current charge
Large fraction of co-propagating dark current at similar energy levels as beam
Halo

## Slide 12: Dark current studies with defects
Gunlab and defect analyzed
Beamtube HOM
0.75/43.0
Bellow
22.3/6.3
Taper
13.4/5.5
Beamtube Steel
50.4/33.0
Full-cell
1.9/1.8
Half-cell
5.9/5.4
Cathode Plug
5.0/4.6
E0=30 MV/m
Fractional losses of all surface losses here in % with Solenoid 75 mT
Solenoid off
Losses in Watt at 1.8 K
Cathode Plug*: 8.5
Half-cell: 9.9
Full-cell: 3.3
Limit by 1.8 K cryo budget
and cathode cooling (32% of total)
*actually at 80K, but part of cavity system
Ratio of dark current
level from cavity exit
to 1st screen to 2nd screen: 9:3:1  only about 11% leave to Booster
Emission peak shifted
by 45 deg. from optimum phase (Ekin,max)
Some particles in SEY>1 region of Nb  rise of multipacting

## Slide 13: Intermediate result
By the damage applied by the manufacturer during the last cleaning before shipping Gun 1.1 to HZB we lost about 2 years to develop and perform a SRF gun recovery program, but:
As shown here, Gun 1.0 in the state as produced would have not sufficed for any kind of ERL operation
We are now able to recover cavities with a harmful damage of the surface layer and do the follow- up treatment by chemistry and high-pressure water rinsing specialized for SRF gun structures, also HF rinsing possible to remove eventual oxidizations
In vertical tests before cleanroom assembly, both Gun cavities were without any field emission
We better preserve this state

## Slide 14: Impressions of Gun cold string assembly
Recent activities
Petrov+cooler positioning
Determine cathode position
Concentricity check with dummy cathode
Completed cold string leak tight and evacuated, accomplished for Nikolaus 2021!
Preassembled sub- components
Detailed, accurate preparation led to quick, but not imprecise assembly
Key was documentation of the past (2016) and a motivated and well working team
Thanks to: André, Michael, Henry, Jan, Frank and fast support from the workshop

## Slide 15: Cold String Role out and preparation for cold string test in HTA
Recent activities
Prior to module assembly, a full cold string test is the final acceptance test in HoBiCaT:
Test if Q0 and max. E-field is preserved after assembly
Did we activate any field emitters by improper handling?
Test of tuner system, coarse and piezo stacks
Characterization of parameters for LLRF control
Cold String inside HoBiCaT
Cold String test assembly
Role out

## Slide 16: Cold string test:First power ramp
Power ramp
At 14:00, 26th of January, the measurements were stopped
by an event:
Vacuum activity measured by both beamline sensors
Losses increase from static level to 2.2 g/s
No radiation was observed downstream
This happened at about E0=19.6 MV/m, known as MP prone regime due to shortened geometry of half cell
Note: no heating of beam tube observed
LF detuning -3.6 Hz/(MV/m)² confirmed field level
A second power ramp showed strong radiation, aka field emission, the test was stopped, RF processing did not work “Hard” Field emitter activated at 19.6 MV/m

## Slide 17: Current SRF gun activities
Current Status
Cathode transfer system in cleanroom:
Optimize transfer mechanics and insert placement into cathode cooler/RF filter unit
Training of transfer procedure, monitor mechanical feedback
Prepare transfer system for in-vacuum particulate counter test
Gun cold string being disassembled
Assembly training with “used” SRF gun parts
Assembly of cold string with second gun and newly conditioned power couplers (CW optimized XFEL/TTF-III version)
Cold mass installation
Cryo module and cold mass integration
Setup and alignment of cryo-module in accelerator hall, cryo- and vacuum connections, diagnostics, RF, LLRF
Vacuum and cooldown tests
RF test without cathode
First cathode transfer (metal plug only)
RF test with cathode, adapting cathode position
Cathode transfer, Laser on cathode, beam

## Slide 18: Status of the Booster
SRF gun
3x2cell Cornell style cavities
Modified coupler section to house a redesigned version of the KEK CERL couplers
120 kW travelling wave per coupler at 100 mA
For first stage, loaded Q is increased to be optimized for max. current of SRF gun
Coupler tip optimized for improved coupling, low penetration depth, low kick and emittance dilution
Cornell style SiC beamtube HOM absorber
Zero-crossing
2.1 MV
2.1 MV

## Slide 19: High power Booster module parts
Booster
Cavities outperform specs
Coupler are being conditioned
One cavity has been tested with tuner and magnetic shielding
HOM arrival last winter all parts for cold string in house
Xelera Research
HOM absorber
parts

## Slide 20: High power coupler conditioning for Booster
Booster fpc
A lot of diagnostics and
Cooling water manifold
Assembly stages

## Slide 21: Parameters under observation
Booster fpc
TW
SW
Max. field localized 10.6 A/m
Reflecting end
Max. field at transition, fluctuates between FPCs 5.7 A/m
CST FD Solver: Pin=0.5 W rms
TW electric field pattern
Field amplitude x2
Phase shift
By coupled RF- thermal simulations potential hotspots were identified and equipped with temperature sensors
All cooling water channels kept separated to learn, where most of the load will happen

## Slide 22: First trials
Booster fpc
95 kW TW reached at 10% duty cycle, vacuum processing starts with 70 kW
Improved bake out procedure and replaced CF gaskets with RF sealed version, improved S11 parameter by 25 dB
Orange: CF, Green, parts covered
by new gasket type

## Slide 23: Follow-up results
Booster fpc
Ca 0.25 K/kW at ceramic window about 55 °C at 120 kW TW expected
Confirms the values obtained in simulations
Imbalance in couplers?
LOS of sensors?
Pulses from 1‰ to 980‰
1e2

## Slide 24: Achievements so far  High Power Coupler Processing
Booster fpc
Parameter | Required Gun 1: 6 mA | Required Gun 2: 100 mA | Values reached so far
Traveling wave | 7 kW | 110 kW | 60 kW CW 120 kW 10% Duty Cycle
Standing wave | 1.75 kW | 54 kW | 15 kW
3 out of 4 pairs conditioned, 4th pair in winter 2022/23

## Slide 25: Summary
And?
Currently, we focus on finalizing the SRF gun, which means to mainly exchange the cavity and couplers
The SRF gun will be installed in the injector and deliver first beam
The booster is going to be assembled, once the gun is installed (not enough personnel for parallel work)
This winter, the last, currently only back-up coupler pair, will be processed to full level: 120 kW TW and 30 kW SW (with overhead)
We better come back next ERL workshop with beam results
Thanks for your attention and to all co-workers at HZB and our collaboration partners,
who supported us on this way

## Slide 26: Module layout
Achievements
80K helium
shield
2nd magnetic
shield
Holders and bellows for alignment
External Solenoid hexapod mover with feedthroughs
Cathode transfer system port
Beam
HGRP
4K cooling:
  Solenoid
  HOM load
  FPCs
4K filling line cavity
1.8 K JT line cavity
80K FPC and HOM
80 K shield and cathode
Both doors can be opened
Solenoid shielding:
Cryoperm around solenoid might saturate
Replaced by Nb disc between Solenoid and cavity outer shield, efficiency factor 5-8 to be published by J. Völker et al.
Modified: Issue with thermal short
Module layout of bERLinPro SRF Gun

## Slide 27: Transfer system and cathode carrier (HZDR modification)
Achievements
Module+transfer system
Transfer system behind module
Cathode cooler (80 K Helium) and Petrov filter holding the insert with tuner
Cathode insert with Plug
and transfer rod
3(4) thermal transitions
Plug: Ø 10 mm

## Slide 28
Major project challenge: 3 different type of cavities and modules for different requirements