## Slide 1: V11 Test Procedures Comparison
Xuan Chen

## Slide 2: Procedures
1) IV curve
2) pixel alive
2a) Noise
3) thradj @3500 e-
4) threqu
5) Scurve
6) [Noise]
7) thradj @2000
8) Scurve
9) [Noise]
10)thradj @1500
11) [Scurve]
12)Noise
13) Scurve
Method 1 (Current benchmark) :
*steps in [ ] were skipped
Method 2 :
1) IV curve
2) pixel alive
2a) Noise
3) thradj @3500 e-
4) threqu
5) Scurve
6) thradj @2000
7) threqu
8) Scurve
9) thradj @1500
10) threqu
12)Noise
13) Scurve

## Slide 3: Procedures
Managed to trim down to 1000 e- in both case
Files:
  /nfs/cms/fpixmech/module_tests/data/cmspixel/runDir_server/xuan/trim_f5_v11_w_threq_0406/
  /nfs/cms/fpixmech/module_tests/data/cmspixel/runDir_server/xuan/trim_f5_v11_wo_threq_0405/

## Slide 4: Results – 3500 e-

## Slide 5: Results – 3500 e-

## Slide 6: Results – 2000 e-

## Slide 7: Results – 1500 e-

## Slide 8: Results – 1200 e-

## Slide 9: Results – 1000 e-

## Slide 10: Results
With threq the thresholds seem to be significantly better
However, the number of masked pixels also increased significantly
Masked Pixels | Without threq | With threq
1500 e- | 0 | 11
1200 e- | 11 | 141
1000 e- | 139 | 533
Chip 7:
Masked Pixels | Without threq | With threq
1500 e- | 0 | 6
1200 e- | 3 | 130
1000 e- | 125 | 531
Masked Pixels | Without threq | With threq
1500 e- | 0 | 0
1200 e- | 14 | 84
1000 e- | 91 | 674
Masked Pixels | Without threq | With threq
1500 e- | 4 | 76
1200 e- | 41 | 343
1000 e- | 693 | 1360
Chip 6:
Chip 5:
Chip 4: