## Slide 1: Beams for medical treatment
Massimo Dal Forno
July 21, 2023
2023 CBB Annual Symposium: Accelerators and Society

## Slide 2: Outline
Introduction
Radiation therapy
Imaging techniques
ViewRay MRIdian Linac
                Magnet, rf system, MLC
Clinical applications

## Slide 3: Radiation therapy, Bragg peak
Types
Electrons
Photons
Protons
Carbon Ions
Beam
entrance
Beam
entrance

## Slide 4: Treatment with photons (X-rays, γ-rays)
Where:

## Slide 5: Radiation therapy
Deliver radiation to the tumor, sparing surrounding organs
“Shoot” from several angles (machine has a gantry that rotates around the patient)
Radiation sums up at isocenter
Courtesy U Wisconsin
Isodose contours

## Slide 6: Seeing, It’s the unfair advantage
Surgeons cut open patients and see all of the motion and treat around it
Radiation oncologist need to see the motion in real time

## Slide 7: Imaging techniques
CT (Computerized Tomography)
MRI (Magnetic Resonance Imaging)
Bladder
Prostate
Rectum
Courtesy Wash U
No radiation delivered when imaging!
ViewRay integrates Linac with MRI

## Slide 8: ViewRay Inc. established in 2004
FDA Cobalt 60 - May 2012, first treatment January 2014
FDA Linac - February 2017, first treatment July 2017

## Slide 9: The ViewRay paradigm
*
* We can only win at the intersection
don’t miss the target

## Slide 10
Two unique technologies
  A 0.35T MRI magnet for imaging the patient
  A linac for Radiation Therapy
Optimized Magnet
  Clear, high SNR images
  Continuous fast imaging without overheating patient
  Geometric precision with high spatial integrity
Radiation source silent to the MRI
  Accurate dose delivered as planned
  No image interference, no image artifacts
Compact system fits in existing rooms
Protected by US and international patents
Diagnostic magnets are optimized for diagnosing patients
ViewRay’s magnet is optimized for treating patients
ViewRay MRIdian Linac
Magnetic & RF
shielding technology*
Linear
Accelerator*
Split MRI magnet
Double-focused
MLC*
Patient table

## Slide 11: LINAC Based Radiation System
Compact inline S-Band 6 MV standing wave linac with side coupled cavities;
Designed for >6 Gy/min
±180º gantry rotation
  360º treatment around patient;
Source-Axis distance = 90 cm (vs 143.5 cm of competitive systems);
Triode e-gun, designed for beam gating (competitive systems don’t have this).
S-Band
Linac
Circulator
Magnetron
Pulse
Transformer
Gun
Driver

## Slide 12: Gantry

## Slide 13: Split magnet

## Slide 14: ViewRay Linac installed in the vault

## Slide 15: Shielding the magnetic field
Concentric
layers of steel

## Slide 16: Building concentric ferromagnetic rings we create voids in the magnetic field to shield the Linac components
Magnetic field lines
Similar to Russian nesting dolls
Magnetic
free zone
All linac compact on gantry & shielded

## Slide 17: Magnetic Shield Technology

## Slide 18: B Field Plots in the Bucket
Magnetic field is reduced at Gauss level

## Slide 19: RF Cloaking Technology

## Slide 21: Double Stacked Double Focused Multi Leaves Collimators (MLC)
138 leaf Double-focus, Double-stack MLC
  68 leaf upper stack; 70 leaf lower stack
  Upper stack offset by ½ leaf width
  MLC faces stays parallel
  Minimum penumbra
Sharpest radiosurgery beam!
Notes: • Rotating Gantry Assembly
• 3 Independent Co60 Headed Design
• Enabling Fast IMRT, SBRT or Conformal External Beam Radiation with Asynchronous Delivery
• Mounted with 120 degree separation
• 15,000 Ci per source
• ±240 degree Rotation for 2 or 3 Head Operation for Maximum Reliability.
• 3 Doubly Focused MLC Systems
• 180 MLC Leaves. 60 per Head
• Best-in-class MLC for Reduced Penumbra & Interleaf Leakage

## Slide 22: Double Stacked Double Focused MLC

## Slide 23: Double Stacked Double Focused MLC

## Slide 24: Clinical Applications

## Slide 25: Gating with Organ Motion
Pink: User define contour: The Radiation beam is shot inside this contour
Red: Contour of the organ to treat (obtained with organ MRI-imaging and auto-contouring)
  Gating mode:
  when the organ moves out of the contour:
    The MRI computer vision sees that the organ is moving out of the contour;
  The machine turns off the radiation beam
Notes: From Left to right
Box 1 – Pancreas
Box 2 – Lung
Box 3 – Prostate
Box4 – Breast

## Slide 26: Day by day variations of organ distances
mm are miles!
ViewRay can see these changes and on-table adapt
Notes: From Left to right
Box 1 – Pancreas
Box 2 – Lung
Box 3 – Prostate
Box4 – Breast

## Slide 27: Day by day variations of organ position
Day 1
Day 2
Sensible organ
Courtesy VUmc
Organ to be treated
Notes: From Left to right
Box 1 – Pancreas
Box 2 – Lung
Box 3 – Prostate
Box4 – Breast

## Slide 28: Day by day variations of organ position
Allows accounting for day to day changes
On-table adaptive radiation therapy (oART)
Notes: From Left to right
Box 1 – Pancreas
Box 2 – Lung
Box 3 – Prostate
Box4 – Breast

## Slide 29
Right Breast
Mediastinum
Head/Neck
Right Lung
Bladder
Mesenteric
Aortic Arch
Hilum
Pancreas
Mediastinal LN
Adrenal
Rib
Pelvic LN
Paraaortic
Liver
Iliac
Stomach
Periaortic
Soft Tissue Neck
Spleen
Spine
Colon
Left Breast
Prostate
Soft Tissue Pelvis
Rectum
Soft Tissue ABD
Left Lung