## Slide 1: Designing and Commissioning an EDX System
Photoemission and Bright Beams Lab

## Slide 2: The Growth Chamber, PBBL

## Slide 3: Energy Dispersive X-Ray Spectroscopy
Energy Dispersive X-Ray Spectroscopy can help us chemically characterize a sample
An electron beam incident on the surface of a sample leads to characteristic X-ray emission along with secondary electrons being emitted

## Slide 4: Designing an EDX System for the Growth Chamber
The electron beam is supplied by a RHEED 15kV electron gun
Add clearing electrodes to deflect secondary electrons
Incident electrons from gun
X-rays and secondary electrons
Sample
X-ray Detector
Clearing electrodes

## Slide 5: EDX Setup Design
Based on the energy distribution of secondary electrons, most have a kinetic energy < 1 keV
Taking this as an upper limit, a relation can be derived between electrode length and required potential difference
For d = 2cm, an electrode length of 4cm was chosen requiring a 1kV supply
Working out the Length of Clearing Electrodes
grounded

## Slide 6: EDX Setup Design
Using caliper measurements, a rough model of the detector was designed for modelling purposes
Modelling the Detector - Cross Assembly
Inventor model for detector

## Slide 7: EDX Setup Design
The protruding piece up front is meant to hold the detector and screw it down there
The second cut on the top is space to hold the positive electrode and screw it down
The ground electrode will be wedged in between the detector and its holder piece, grounding the whole setup
Modelling the Modified Circular Rod
Front view of the Modified Circular Rod

## Slide 8: EDX Setup Design
Hit separation distance between the electrodes
The washers electrically isolate the screws to avoid shorting
The connector (also isolated) for the positive electrode is wedged between the washers and the electrode
The back end is also modified to lock down the orientation of the rod for welding
Modelling the Modified Circular Rod
Back view of the Modified Circular Rod
View of the Detector Rod Assembly Under the Cross

## Slide 9: EDX Setup Design
The holes for the screws holding the positive electrode go all the way through the rod to prevent any trapped air
Matching the end of the rod, the blank flange is also modified to have the rod sit in neatly for the welding
Features of the Assembly Model
View of the Detector with Clearing Electrodes
Modified Blank Flange

## Slide 10: EDX Assembly
Minor adjustments made regarding spacing washers and translator setup
Due to less clearance, the wires were bound about the rod
Assembling the Cross
Detector Cross and Electron Gun
Detector in Growth Chamber

## Slide 11: EDX Assembly
The raw data collected by the SDD needs to be converted to readable data
Amptek’s DppMca is used to acquire spectra and control the detector and processor setup
PyMca is used for advanced qualitative and quantitative analysis
Assembling the Detector Setup
Preamplifier
Processor

## Slide 12: DPPMCA
For a sample, we observed spectra with the high voltage supply turned on and off
Clearer peaks, reduced Bremsstrahlung curve
Impact of Clearing Electrodes

## Slide 13: DPPMCA
Incorrect fits on P47 scintillator screens and Mo sample
Issues with setting the noise thresholds without signal
Working on fixes
Troubleshooting

## Slide 14: Acknowledgements
Dr. Siddharth Karkare
Peter Owusu
Photoemission and Bright Beams Lab
Center for Bright Beams
Thank You! Any Questions?